## Sheet1
2017-11-30 BS OL Cal |  |  | Mark Barton
Calibration and Diagonalization of BS TILT OL
Calibration of QPD
Horizontal Movement of QPD |  |  | Data by Shoda-san and Mark
Micrometer | Zeroed | QPD*100 | QPD
8.14 | -0.435 | 55 | 0.55
8.24 | -0.335 | 48 | 0.48
8.34 | -0.235 | 37 | 0.37
8.44 | -0.135 | 22 | 0.22
8.54 | -0.035 | 5.5 | 0.055
8.56 | -0.015 | 2 | 0.02
8.575 | 0 | -0.5 | -0.005
8.575 | 0 | 0 | 0
8.64 | 0.065 | -11 | -0.11
8.74 | 0.165 | -26 | -0.26
8.84 | 0.265 | -38 | -0.38
8.94 | 0.365 | -48 | -0.48
9.04 | 0.465 | -54 | -0.54
9.14 | 0.565 | -58 | -0.58
Slope: |  |  | -1.5281 | cts/mm
Vertical Movement of QPD |  |  | Data by Shoda-san and Mark
Micrometer | Zeroed | QPD*100 | QPD
4.32 | -0.5 | -56 | -0.56
4.42 | -0.4 | -50 | -0.5
4.52 | -0.3 | -42 | -0.42
4.62 | -0.2 | -28 | -0.28
4.67 | -0.15 | -22 | -0.22
4.72 | -0.1 | -15 | -0.15
4.77 | -0.05 | -8 | -0.08
4.82 | 0 | 0 | 0
4.92 | 0.1 | 17 | 0.17
4.97 | 0.15 | 23 | 0.23
5.02 | 0.2 | 30 | 0.3
5.07 | 0.25 | 36 | 0.36
5.12 | 0.3 | 41 | 0.41
5.17 | 0.35 | 45 | 0.45
5.22 | 0.4 | 49 | 0.49
5.27 | 0.45 | 52 | 0.52
Slope: |  |  | 1.4535 | cts/mm
Response to Torques at IM |  |  | Data by Mark
Yaw Torque at IM
# | IM_TEST_Y | QPDY*100 | QPDP*100 | IM_DAMP_Y | IM_DAMP_P | QPDY | QPDP
12 | -200 | 46 | -58 | -356.5 | -57.3 | 0.46 | -0.58
11 | -160 | 41.6 | -59 | -325 | -57.35 | 0.416 | -0.59
10 | -120 | 35 | -55.5 | -293 | -57.4 | 0.35 | -0.555
9 | -80 | 25.5 | -46.5 | -261 | -57.45 | 0.255 | -0.465
8 | -40 | 14.5 | -31.5 | -228.5 | -57.55 | 0.145 | -0.315
1 | 0 | 3 | -11.5 | -196 | -56.95 | 0.03 | -0.115
7 | 0 | 3 | -12 | -196 | -57.55 | 0.03 | -0.12
2 | 40 | -8.5 | 10.5 | -163 | -57 | -0.085 | 0.105
3 | 80 | -19.5 | 31 | -129.5 | -57.1 | -0.195 | 0.31
4 | 120 | -29 | 46 | -95.5 | -57.45 | -0.29 | 0.46
5 | 160 | -37.5 | 55.5 | -61.5 | -57.9 | -0.375 | 0.555
6 | 200 | -45 | 60.25 | -26.5 | -58.15 | -0.45 | 0.6025
Slopes: |  |  |  | 0.82386 | -0.0012159 | -0.0025753 | 0.0048125
|  |  |  | µrad/ct(A) | µrad/ct(A) | ct/ct(A) | ct/ct(A)
Calibrated, diagonalized QPD data
[6 empty] | QPDY | QPDP
[6 empty] | -243.7181 | 38.3813
[6 empty] | -239.8424 | 26.2722
[6 empty] | -219.1869 | 14.5517
[6 empty] | -177.6944 | 2.7872
[6 empty] | -116.0585 | -4.9342
[6 empty] | -38.8028 | -7.4419
[6 empty] | -40.287 | -8.0863
[6 empty] | 44.3896 | -7.3723
[6 empty] | 122.3516 | -8.0061
[6 empty] | 181.6544 | -12.0387
[6 empty] | 223.0756 | -20.6997
[6 empty] | 248.8415 | -33.0227
[5 empty] | Mysterious Calibration Factor Difference: | 2.1713
Pitch Torque at IM
# | IM_TEST_P | QPDY*100 | QPDP*100 | IM_DAMP_Y | IM_DAMP_P | QPDY | QPDP
12 | -1000 | -63 | -47 | -194.8 | -71.4 | -0.63 | -0.47
11 | -800 | -53 | -41 | -195 | -69 | -0.53 | -0.41
10 | -600 | -41 | -36 | -195 | -66.65 | -0.41 | -0.36
9 | -400 | -27.5 | -31 | -195.5 | -64.3 | -0.275 | -0.31
8 | -200 | -7 | -15 | -195.5 | -61.15 | -0.07 | -0.15
1 | 0 | 3 | -11.5 | -196 | -58.35 | 0.03 | -0.115
7 | 0 | 3 | -11 | -196 | -58.8 | 0.03 | -0.11
2 | 200 | 13 | -7 | -196 | -56.2 | 0.13 | -0.07
3 | 400 | 22.5 | -3 | -196.5 | -54.1 | 0.225 | -0.03
4 | 600 | 31.5 | 1 | -196 | -51.95 | 0.315 | 0.01
5 | 800 | 39 | 5 | -196.5 | -49.8 | 0.39 | 0.05
6 | 1000 | 46 | 9.5 | -196 | -47.6 | 0.46 | 0.095
Slopes: |  |  |  | -0.00079545 | 0.012057 | 0.00056841 | 0.00029159
|  |  |  | µrad/ct(A) | µrad/ct(A) | ct/ct(A) | ct/ct(A)
Calibrated, diagonalized QPD data
[6 empty] | QPDY | QPDP
[6 empty] | -41.5192 | -215.4969
[6 empty] | -39.2636 | -183.1729
[6 empty] | -43.0874 | -147.2191
[6 empty] | -49.2444 | -107.5765
[6 empty] | -33.6374 | -36.5443
[6 empty] | -38.8028 | -7.4419
[6 empty] | -37.3186 | -6.7976
[6 empty] | -40.9998 | 22.9491
[6 empty] | -43.9033 | 51.4662
[6 empty] | -46.029 | 78.7536
[6 empty] | -45.8216 | 102.3523
[6 empty] | -43.3521 | 125.3657
[6 empty] | Mysterious Calibration Factor Difference: | 14.322
Analysis
Lever Arm Calc |  |  |  |  | QPD Calibration
Internal to tank |  |  | 976 | 2017-11-07 | "Yaw" | "Pitch"
Viewport glass to beam spot on steering mirror: 117 mm. |  |  | 105 | 2017-11-07 | 335.1239 | 277.6388 | µrad/ct
Steering mirror to beamsplitter: 28 mm (rough; estimated from photo). |  |  | 34 | 2017-11-07
Beamsplitter to TILT QPD |  |  | 124 | 2017-12-01
Total lever arm: |  |  | 1239
|  | rad | °
Beam angle |  | 0.66323 | 38 | 2017-12-05
Matrix from IM Torque [Y;P] to IM Angle [Y;P] - µrad/ct(A)
[5 empty] | rad | °
| [Y;P]= | 0.82386 | -0.00079545 | .[Y;P] | -0.0014759 | -0.084561
|  | -0.0012159 | 0.012057 |  | -0.06588 | -3.7746
Matrix from IM Torque [Y;P] to Raw QPD ["Yaw";"Pitch"] - kct/ct(A)
[5 empty] | rad | °
| ["Yaw";"Pitch"]= | -0.0025753 | 0.00056841 | .[Y;P] | 2.0622 | 118.1527 | 61.8473
|  | 0.0048125 | 0.00029159 |  | 1.0968 | 62.8424
Matrix from IM Torque [Y;P] to Raw QPD ["Yaw";"Pitch"] - µrad/ct(A)
[5 empty] | rad | °
| ["Yaw";"Pitch"]= | -0.0025753 | 0.00056841 | .[Y;P] | 1.0794 | 61.8473
|  | 0.0048125 | 0.00029159 |  | 1.0968 | 62.8424
|  |  |  | Average: | 1.0881 | 62.3448
Diagonalization and Calibration
|  | rad | °
Rotation Angle |  | 1.0881 | 62.3448
Rotation Matrix
| [Yunc;Punc]= | -0.46415 | 0.88576 | .["Y";"P"]
|  | 0.88576 | 0.46415
Calibration Matrix
| [Y;P]= | 335.1239 | 0 | .[Yunc;Punc]
|  | 0 | 277.6388
Flipped form for typing into lower left corner of screen
Net Calibration*Rotation Matrix |  |  |  |  | VIS_CUST_TYPEB_TM_OPLEV_OL2EUL_MTRX.adl
| [Y;P]= | -155.5474 | 296.8384 | .[Y;P] | [P;Y]= | 128.8657 | 245.9206 | .[TILT_P;TILT_Y]
|  | 245.9206 | 128.8657 |  |  | 296.8384 | -155.5474
For use if Simon's gains of 100 are kept in screen
Net Calibration*Rotation Matrix/100 |  |  |  |  | VIS_CUST_TYPEB_TM_OPLEV_CHAIN.adl
| [Y;P]= | -1.5555 | 2.9684 | .[Y;P] | [P;Y]= | 1.2887 | 2.4592 | .[TILT_P;TILT_Y]
|  | 2.4592 | 1.2887 |  |  | 2.9684 | -1.5555
cf:
For use if Simon's gains of 100 are kept in
Note on C8: MB: As measured after Simon's x100 gain, at K1:VIS_BS_TM_OPLEV_TILT_YAW_OUTMON
Note on D8: MB: With x100 gain removed, as if measured at K1:VIS_BS_TM_OPLEV_TILT_YAW_INMON
Note on C26: MB: As measured after Simon's x100 gain, at K1:VIS_BS_TM_OPLEV_TILT_PIT_OUTMON
Note on D26: MB: With x100 gain removed, as if measured at K1:VIS_BS_TM_OPLEV_TILT_PIT_INMON
Note on C47: MB: As measured after Simon's x100 gain, at K1:VIS_BS_TM_OPLEV_TILT_YAW_OUTMON
Note on D47: MB: As measured after Simon's x100 gain, at K1:VIS_BS_TM_OPLEV_TILT_PIT_OUTMON
Note on E61: MB: The IM OSEMs are believed to be correctly calibrated and diagonalized so that the inputs to the damping filters is in µm or µrad. Thus the slope should be in µrad per count of actuation (currently uncalibrated).
Note on F61: MB: The IM OSEMs are believed to be correctly calibrated and diagonalized so that the inputs to the damping filters is in µm or µrad. Thus the slope should be in µrad per count of actuation (currently uncalibrated).
Note on G77: MB: The apparent difference in yaw calibration factor estimated from the range of the processed QPD yaw data over the range of the OSEM yaw data.
Note on H110: MB: The apparent difference in calibration factor estimated from the range of the processed QPD pitch data over the range of the OSEM pitch data.
Note on F114: MB: We throw away the sign here because we really don't care about the direction relative to the micrometer screw. We include a factor of cos(37°) per klog 3546 and Akutsu-san's email to Mark et al. of 2017-12-05.